Compliance Data Record ISO 17664-2 (2021) Azurion HW R3.

Table of Contents

1	Record history	4

2	Purpose	4

3	Scope	4

4	References	4

5	Definitions & abbreviations	5

6	Compliance Checklist	6

Pre-created CDR history	18

# Record history

| Revision | Release Date | Author | Description of changes | CR / Reason |
| --- | --- | --- | --- | --- |
| A | 2023 Jul 04 | Wim van Rooij | Initial version | N/A |
| B | 2024 Dec 10 | Karel Rysman | Update to include SW R4.0 and R4.0 IFU to Azurion HW R3 | N/A |
| C | 2025 March 1 | Ashwini Lingappa Swamy | Update to include SW R3.1 and IFU R3.1 to Azurion HW R3 Update to IFU addendum Footswitch Update to IFU addendum for wireless footswitch | N/A |
| D | 2026 Mrt 04 | Camila Ávila | Update IFU R4.0 Update the document revision history Update definitions & abbreviations list. Update reference list. | N/A |

# Purpose

This document shows the compliance to ISO 17664-2:2021 [ISO 17664-2] for the Azurion HW R3.

# Scope

Organization name: Philips Medical Systems Nederland B.V. – IGT-S.

This document is meant as compliance evidence for internal use only, but it may be used as evidence for external authorities (e.g., Notified Body).

# References

| Reference Number | Document Title | Document ID |
| --- | --- | --- |
| [IFU- Foot Switch] | IFU Addendum Foot Switch | 4523 001 2578X1 |
| [IFU- LegSupportBoard] | Addendum Leg Support Board 1.0 | 4523 001 1130X1 |
| [IFU- WFS] | IFU Addendum Wireless Foot Switch | 4523 001 2182X1 |
| [IFU-Leaflet] | Leaflet. Philips Azurion - Accessing the Instructions for Use in Electronic Form | 4523 001 1121X1 |
| [IFU] | Instructions for Use Azurion R3.0 Instructions for Use Azurion R3.1 Instructions for Use Azurion R4.0 | 4523 001 0357X1 4523 001 2918X1 4523 001 2500X1 |
| [ISO 17664-2] | Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices –Part 2: Non-critical medical devices | ISO 17664-2:2021(Ed 1.0) |
| [OD-2020] | IECEE OPERATIONAL DOCUMENT TRF – Development, maintenance and use | IECEE OD-2020 |
| [PMSRAR-ISO-17664] | Pre-market Safety Risk Assessment Report ISO 17664:2017 Risk management file impact analysis - ISO 17664-2021 Compliance | DHF332364 Rev.00 D002077627 Rev.A |
| [SCDV] | Summary of cleaning and disinfectant validation | D001417495 Rev.A |
| [Test Protocol] | Test Protocol Record ISO 17664-2 (2021) | D001024021 Rev.C |

- Where x may be 1,2,…,9

# Definitions & abbreviations

| Terminology & Abbreviations | Description/Definition |
| --- | --- |
| CDR | Compliance Data Record |
| DHF | Design History File |
| DMR | Device Master Record |
| DUT | Device Under Test |
| HW | HardWare |
| Hz | Hertz |
| IECEE | IEC System for Conformity Assessment Schemes for Electrotechnical Equipment and Components. |
| IFU | Instructions for Use |
| IGT-S | Image Guided Therapy - Systems |
| ISO | International Organization for Standardization |
| kVA MAX | Maximum kilovolt‑ampere rating |
| N/A | Not applicable |
| NCB | National Certification Body |
| OD | Operational Document |
| PE | Protective Earth |
| REF | REFerence |
| Rx.x | Release x.x |
| TRF | Test Report Form |
| VAC | Volts Alternating Current |

# Compliance Checklist

| ISO 17664-2 Processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices – Part 2: Non-critical medical devices |  |  |
| --- | --- | --- |
| Test specification: |  |  |
| Standard : | ISO 17664-2:2021 (Ed. 1.0) |  |
| Test item description : | Medical diagnostic X-ray systems |  |
| Trade Mark : | Philips |  |
| Manufacturer : | Philips Medical Systems Nederland B.V. |  |
| Model/Type reference : |  |  |
| Ratings : |  |  |
| Testing location : | Philips Medical Systems Nederland B.V. Veenpluis 6 5684 PC, Best, The Netherlands |  |
| Possible test case verdicts: |  |  |
| - test case does not apply to the test object : |  | N/A |
| - test object does meet the requirement : |  | P (Pass) |
| - test object does not meet the requirement : |  | F (Fail) |
| General remarks: |  |  |
| The test results presented in this report relate only to the object tested. Throughout this report a comma / point is used as the decimal separator. |  |  |
| Copy of marking plate: The artwork below may be only a draft. <The use of certification marks on a product must be authorized by the respective NCBs that own these marks> |  |  |
| General product information: <Include a copy of the intended use from the IFU> |  |  |
| The product fulfils the requirements of ISO 17664-2:2021 |  |  |

| Date | 2023-Jul-04 |
| --- | --- |
| Test Location | Philips Medical Systems Nederland B.V. Veenpluis 6 5684 PC, Best, The Netherlands |
| Test Method Standard | ISO 17664-2:2021 (Ed. 1.0) processing of health care products - Information to be provided by the medical device manufacturer for the processing of medical devices – Part 2: Non-critical medical devices [Test Protocol] |
| Test item description | Angiographic X-ray system family |
| Trade mark | Philips |
| Manufacturer | Philips Medical Systems Nederland B.V. |
| Sample Type (Identification of Design) | Azurion series, release: HW R3, model 001501 with the following configurations: Azurion 3 M12 [722 229]1) Azurion 3 M15 [722 230]1) Azurion 5 M12 [722 231]2) Azurion 5 M20 [722 232]2) Azurion 7 M12 [722 233]2) Azurion 7 M20 [722 234]2) Azurion 7 B12 [722 235]1) Azurion 7 B20 [722 236]1) Azurion FlexArm [722 220]1) Software release: Azurion R3.0, R3.1 and Azurion R4.0 1) configuration with Azurion R3.0, R3.1 software 2) configuration with Azurion R3.0, R3.1, R4.0 software |
| Ratings | 3-Phase: 380/400/415/440/480 VAC, 80/76/73/69/63 A, 50/60 Hz, 3 ~ + PE, Momentary power 100 kVA MAX |
| Test Set-up /Test SW | According to the standard |
| Tester | Wim van Rooij |
| Possible test case verdicts: - test case does not apply to the test object : N/A - test object does meet the requirement : P (Pass) - test object does not meet the requirement : F (Fail) |  |
| General remarks: The test results presented in this report relate only to the object tested. Throughout this report a comma / point is used as the decimal separator. |  |
| Copy of marking plate: The artwork below may be only a draft. The use of certification marks on a product must be authorized by the respective NCBs that own these marks. Azurion series, release: HW R3 Device name; REF. Azurion 3 M12; 722 229. Azurion 3 M15; 722 230. Azurion 5 M12; 722 231. Azurion 5 M20; 722 232. Azurion 7 M12; 722 233. Azurion 7 M20; 722 234. Azurion 7 B12; 722 235. Azurion 7 B20; 722 236. Azurion FlexArm; 722 220.. Copy of rating plate |
| General product information: The Azurion HW R3 series consists of: Azurion 3 M12, Azurion 3 M15, Azurion 5 M12, Azurion 5 M20, Azurion 7 M12, Azurion 7 M20, Azurion 7 B12, Azurion 7 B20 and Azurion FlexArm. Product Description The Azurion series consists of a number of monoplane and biplane systems with different detector sizes (12”, 15” and 20”). Indications for Use / Medical Purpose The Azurion series (within the limits of the used operating room table) is intended for use to perform: - Image guidance in diagnostic, interventional, and minimally invasive surgery procedures for the following clinical application areas: vascular, non-vascular, cardiovascular, and neuro procedures. - Cardiac imaging applications including diagnostics, interventional and minimally invasive surgery procedures. Additionally: - The Azurion series can be used in a hybrid operating room. - The Azurion series contains a number of features to support a flexible and patient-centric procedural workflow. The Azurion series is intended for all human patients of all ages. Patient weight is limited to the specification of the patient table. Clinical benefits With Azurion, appropriate patient care can be provided by enabling diagnosis and minimally invasive treatment of vascular, non-vascular, cardiovascular, and neurovascular conditions in all types of patients. When configured with ClarityIQ (option), Azurion provides imaging and image guidance at reduced patient radiation dose levels, thereby reducing the risk of health problems associated with ionizing radiation. Intended Operator Profile The Azurion series is intended to be used and operated by adequately trained, qualified, and authorized healthcare professionals who have an understanding of the safety information and emergency procedures as defined by local laws and regulations for radiation workers and staff. Clinical Environment The Azurion series is a fixed and stationary system that can be used in a clinical environment fulfilling the local laws and regulations for radiological X-ray systems in sterile and non-sterile environments. General Safety and Effectiveness To facilitate safe and efficacious operation of the system by a trained healthcare professional, instructions for use are provided as part of the device labeling and training is provided during system handover. Contraindications Avoid using the system with patients who are pregnant or who may possibly be pregnant. However, the risk may be outweighed by the benefit of diagnosing or treating a serious condition. It is the responsibility of the personnel operating the system to make the decision. Avoid using the system in case of existing radiation injury (operator or patient). Operating Principle The system uses X-ray generation, detection and image processing for medical imaging, and additionally displays images from other sources (for example ultrasound). The control mechanisms are input devices and controls for, for example, geometry and table movements. The system provides feedback by audible and visual means. Model differences: There is an Azurion 3-series, an Azurion 5-series, an Azurion 7-series and an Azurion FlexArm system. The Azurion 3-series are ‘entry level’ systems, the 5-series are the mid level systems and the 7-series are the high end systems. A 3-series model is always a floor stand with less options, e.g. a spring arm boom with maximum 3 monitors, while for the 7-series all options are available, incl. ceiling mounted C-arm. The ‘M’ in the configuration name stands for “Monoplane”, the “B” is for “Biplane”. The number at the end of the configuration name is the diagonal size of the detector in inch. There is a 12, 15 and 20 inch detector, but not all detectors can be used on all configurations. The Azurion FlexArm is a monoplane system, 20 inch detector, with FlexArm, for China only. |  |
| The product fulfils the requirements of ISO 17664-2:2021 |  |

No measurements taken for this Test Record. Any measurements listed in this Test Record are generated by the referenced test reports.

| 4 | Risk analysis |  | P |
| --- | --- | --- | --- |
|  | The medical device manufacturer shall undertake risk analysis to determine the content and detail of the information to be provided. The risk management undertaken by the manufacturer of the medical device shall conform with ISO 14971. | See Safety Risk Assessment Report ISO 17664 [PMSRAR-ISO-17664] | P |
|  | NOTE 1 Some of the points relevant to processing that any risk analysis can require (but not limited to) include: — the likely points of contact with the user and/or the patient that might allow cross-contamination; — nature and design of the medical device; — nature of the contamination on the medical device; ﻿— intended use; — foreseeable user error and misuse; — user training; — equipment required for processing; — accessories and consumables required for processing; — necessary maintenance of the medical device; — post-market information; — service life; — necessary warnings. The points above can also be of benefit to those validating alternative processes in accordance with the NOTE 2 to 5.2. |  | - |
|  | NOTE 2 Annex C provides information on classification of medical devices, which can assist with any risk analysis process. |  | - |
| 5 | Validation of the processes identified in the information provided by the medical device manufacturer |  | P |
| 5.1 | The medical device manufacturer shall validate each process that is identified in the information supplied with the medical device and demonstrate that each process is suitable for processing of the medical device. This shall include the intended points of contact with the user and/or the patient that could likely lead to cross-contamination. Parts of the medical device that are unlikely to lead to cross contamination may be excluded from validation, based upon the risk analysis described in Clause 4. | See summary of cleaning & disinfection validation [SCDV] | P |
| 5.2 | The medical device manufacturer shall have objective evidence available that validation of the processing procedures has been undertaken to confirm that the specific medical device will be clean and/or disinfected, when processed as directed. | See summary of cleaning & disinfection validation [SCDV] Sterilization is not applicable for Azurion HW R3 | P |
|  | NOTE 1 A worst-case approach, representing those areas of the medical device that are the intended points of contact with the user, the patient or both and where there is opportunity for cross-contamination, can be used. |  | - |
|  | NOTE 2 In addition to the duty of a manufacturer to demonstrate the validity of provided information, national authorities can require the effectiveness of the final process to be verified by the processor. |  | - |
|  | NOTE 3 National authorities can allow or require the use of an alternative process. In such cases they usually require validation of those processes by the processor. |  | - |
| 5.3 | If a manufacturer supplies a number of different medical devices that share common attributes, then validation studies may be performed as a product family. If this approach is taken, the medical device manufacturer shall demonstrate commonality between the different medical devices and the validation studies shall address the worst-case attribute(s) of the product family. | See summary of cleaning & disinfection validation [SCDV] | P |
|  | NOTE For guidance on grouping of medical devices as product families, see C.1. |  | - |
| 6 | Information to be provided by the medical device manufacturer |  | P |
| 6.1 | General |  | P |
| 6.1.1 | The information specified in this clause shall take into account the nature of the medical device, its intended location for use and processing and its intended use. | Considered See: [IFU] [IFU- LegSupportBoard] [IFU- Foot Switch] [IFU- WFS] | P |
|  | NOTE The location of processing can be at the point of use or within a processing department. |  | - |
| 6.1.2 | The medical device manufacturer shall specify validated method(s) to reduce the risk of transmission of infectious agents to a level appropriate for the intended use of the medical device. | See: [IFU] Azurion R3.0 IFU, section 15.2 Azurion R3.1 IFU, section 15.2 Azurion R4.0 IFU, section 15.2 [IFU- LegSupportBoard], section 4.2 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
| 6.1.3 | Medical device manufacturers shall specify in their processing instructions any special techniques and accessories that will enable the processor to provide a medical device that is suitable for its intended use. | See: [IFU] Azurion R3.0 IFU, section 15.2 Azurion R3.1 IFU, section 15.2 Azurion R4.0 IFU, section 15.2 [IFU- LegSupportBoard], section 4.2 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
| 6.1.4 | When providing processing instructions, medical device manufacturers shall be aware of: |  |  |
|  | — available national and international standards and guidelines; | European and North American standards and guidances taken into account. The system is classified as non-critical | P |
|  | — if applicable, the need for specific training; | Cleaning and wipe disinfection are processes which are familiar to the intended users. No special training required | N/A |
|  | — the processing equipment commonly available to the processor. | No special processing equipment is required | N/A |
|  | NOTE 1 Some national standards and regulations require cleaning and disinfection for all non-critical medical devices. |  | - |
|  | NOTE 2 Annex A provides information which can assist with identifying the information required and the processing equipment commonly available. |  | - |
| 6.1.5 | The equipment or materials required in the specified processes shall be identified by their generic names or specification. Trade names may be added in cases where generic names do not provide sufficient information (see Annex D). | Generic names of materials are specified, see: [IFU] - Azurion R3.0 IFU, section 15.2.2 to 15.2.4 - Azurion R3.1 IFU, section 15.2.2 to 15.2.4 - Azurion R4.0 IFU, section 15.2.2 to 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
| 6.2 | Processing instructions |  | P |
| 6.2.1 | At least one validated method shall be specified for each applicable stage of processing of the medical device. Disinfection may be carried out concurrently with cleaning of the medical device. | See: [IFU] - Azurion R3.0 IFU, section 15.2.2 to 15.2.4 - Azurion R3.1 IFU, section 15.2.2 to 15.2.4 - Azurion R4.0 IFU, section 15.2.2 to 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | NOTE1 Annex A provides information on the commonly used processes available. |  | - |
|  | NOTE 2 The requirements for cleaning and disinfection are stated as separate clauses in this document. However, when the steps are concurrent, the requirements of both stages can be considered as one. In such cases removal of soil, a reduction in microorganisms and inactivation of viable microorganisms can be achieved as a result of the combination of applying the disinfecting agent and a physical action. |  | - |
|  | NOTE 3 The range of medical devices included within this document is wide and varied. Many of these medical devices (e.g. stethoscopes and blood pressure cuffs) are relatively simple medical devices which do not necessarily require automated processes. There will be other medical devices where automated processing is not possible or contra-indicated (e.g. some medical devices with electronic components). However, some medical devices such as beds, wheelchairs and footwear can be, and often are, subjected to automated processes. For this final group of medical devices, a validated method of automated processing can be specified and is preferred. |  | - |
| 6.2.2 | The method shall be appropriate to the market in which the medical device is to be supplied. | Considered | P |
| 6.2.3 | The following information shall be stated where it is required for the maintenance of the intended function of the medical device and the safety of the user(s) and the patient: | See: [IFU] - Azurion R3.0 IFU, section 15.2 - Azurion R3.1 IFU, section 15.2 - Azurion R4.0 IFU, section 15.2 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | a) details of process steps; | See above | P |
|  | b) a description of the equipment the accessories or both; | See above | P |
|  | c) specifications for process parameters and, if applicable, their tolerances. | See above | P |
|  | NOTE For an example of an appropriate format see Annex B. |  | - |
| 6.3 | Limitations and restrictions on processing |  | P |
| 6.3.1 | If processing of a medical device in accordance with the medical device manufacturer’s instructions is known to lead to degradation that can limit the service life of the medical device, then the medical device manufacturer shall provide such information regarding limitations and restrictions to the processor. | See: [IFU] - Azurion R3.0 IFU, section 15.2.6 - Azurion R3.1 IFU, section 15.2.6 - Azurion R4.0 IFU, section 15.2.6 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
| 6.3.2 | If the service life of the medical device is limited by the number of processing cycles or some other end of life indicator(s) this information shall also be provided. | See: [IFU] - Azurion R3.0 IFU, section 15.2.6 - Azurion R3.1 IFU, section 15.2.6 - Azurion R4.0 IFU, section 15.2.6 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
| 6.3.3 | Where an incompatibility of the medical device with a commonly used substance(s) or processing condition(s) that can impact upon patient safety is known, this information shall be provided. | See: [IFU] - Azurion R3.0 IFU, section 15.2.5 - Azurion R3.1 IFU, section 15.2.5 - Azurion R4.0 IFU, section 15.2.5 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
| 6.4 | Preparation before processing |  | P |
|  | If preparation of a medical device is required prior to cleaning to ensure effective processing of that medical device, then the following information shall be provided, where applicable: | See: [IFU] - Azurion R3.0 IFU, section 15.2.5 - Azurion R3.1 IFU, section 15.2.2 - Azurion R4.0 IFU, section 15.2.2 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | a) a description of the process for disassembly of the medical device; | Disassembly of the medical device is not required. See: [IFU] - Azurion R3.0 IFU, section 15.2.5 - Azurion R3.1 IFU, section 15.2.2 - Azurion R4.0 IFU, section 15.2.2 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | N/A |
|  | b) a description of the process and processing parameters for medical device preparation; | See: [IFU] - Azurion R3.0 IFU, section 15.2.5 - Azurion R3.1 IFU, section 15.2.2 - Azurion R4.0 IFU, section 15.2.2 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | c) accessories and tools required. | See: [IFU] - Azurion R3.0 IFU, section 15.2.5 - Azurion R3.1 IFU, section 15.2.2 - Azurion R4.0 IFU, section 15.2.2 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | NOTE For detailed guidance see Annex A. |  | - |
| 6.5 | Cleaning |  | P |
| 6.5.1 | General |  | P |
|  | At least one validated cleaning method shall be specified. | See: [IFU] - Azurion R3.0 IFU, section 15.2 - Azurion R3.1 IFU, section 15.2 - Azurion R4.0 IFU, section 15.2 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
| 6.5.2 | Manual cleaning |  | P |
|  | If a manual cleaning method is specified, the following information shall be included, where applicable: | See: [IFU] - Azurion R3.0 IFU, section 15.2.3 - Azurion R3.1 IFU, section 15.2.3 - Azurion R4.0 IFU, section 15.2.3 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | a) a description of the manual method with step-by-step instructions and the sequence of each individual process step; | See: [IFU] - Azurion R3.0 IFU, section 15.2.3 - Azurion R3.1 IFU, section 15.2.3 - Azurion R4.0 IFU, section 15.2.3 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | b) a description of the process and processing parameters and, if applicable, any limits the medical device can withstand; | See: [IFU] - Azurion R3.0 IFU, section 15.2.3 - Azurion R3.1 IFU, section 15.2.3 - Azurion R4.0 IFU, section 15.2.3 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | c) a description of the accessories required; | See: [IFU] - Azurion R3.0 IFU, section 15.2.3 - Azurion R3.1 IFU, section 15.2.3 - Azurion R4.0 IFU, section 15.2.3 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | d) identification and concentration of process chemicals required; | See: [IFU] - Azurion R3.0 IFU, section 15.2.3 - Azurion R3.1 IFU, section 15.2.3 - Azurion R4.0 IFU, section 15.2.3 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | e) the contact time of any cleaning agents used; | The contact time of the cleaning agent with the equipment is not relevant. | N/A |
|  | f) the quality of water to be used; | See: [IFU] - Azurion R3.0 IFU, section 15.2.3 - Azurion R3.1 IFU, section 15.2.3 - Azurion R4.0 IFU, section 15.2.3 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | g) methods to be used for rinsing (including the need for rinsing between cleaning and subsequent steps where the process residues could interact adversely with the disinfecting agent); | See: [IFU] - Azurion R3.0 IFU, section 15.2.3 - Azurion R3.1 IFU, section 15.2.3 - Azurion R4.0 IFU, section 15.2.3 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | h) if known, identification of any incompatibilities of cleaning agents with the medical device. | See: [IFU] - Azurion R3.0 IFU, section 15.2.5 - Azurion R3.1 IFU, section 15.2.5 - Azurion R4.0 IFU, section 15.2.5 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | The medical device manufacturer's instructions for use may direct the processor to refer to the cleaning product manufacturer's instructions for use. | See: [IFU] - Azurion R3.0 IFU, section 15.2 - Azurion R3.1 IFU, section 15.2 - Azurion R4.0 IFU, section 15.2 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
| 6.5.3 | Automated cleaning |  | N/A |
| 6.5.3.1 | If the automated cleaning process recommends the use of a washer-disinfector meeting the requirements of the ISO 15883 series, the information regarding the automated process may be limited to those parameters that are specific for the medical device, such as specific load configuration, positioning, connection, accessories, process chemicals, water quality, pressures or temperature limit(s) and a statement confirming the recommendation to use a washer-disinfector conforming with the ISO 15883 series. | Automated cleaning is not possible | N/A |
| 6.5.3.2 | If the specific cleaning requirements of the medical device do not allow a generic claim of compatibility with a washer-disinfector meeting the requirements of the ISO 15883 series then the following information shall be included, where applicable: | Automated cleaning is not possible | N/A |
|  | a) a description of the process and processing parameters including any limits the medical device can withstand; | Automated cleaning is not possible | N/A |
|  | b) a description of the accessories required; | Automated cleaning is not possible | N/A |
|  | c) identification and concentration of process chemicals required; | Automated cleaning is not possible | N/A |
|  | d) the contact time of any cleaning agents used; | Automated cleaning is not possible | N/A |
|  | e) the quality of water to be used; | Automated cleaning is not possible | N/A |
|  | f) methods to be used for rinsing (including the need for rinsing between cleaning and subsequent steps where the process residues could interact adversely with the disinfecting agent); | Automated cleaning is not possible | N/A |
|  | g) if known, identification of any incompatibilities of cleaning agents with the medical device. | Automated cleaning is not possible | N/A |
|  | The medical device manufacturer's instructions for use may direct the processor to refer to the cleaning product manufacturer's instructions for use. | Automated cleaning is not possible | N/A |
|  | NOTE Additional information conforming with 6.5.3.2 can be provided when the requirements of 6.5.3.1 are met if the medical device manufacturer chooses to do so. |  | - |
| 6.6 | Disinfection |  | P |
| 6.6.1 | General |  | P |
|  | If disinfection is required, at least one validated disinfection method shall be specified. | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | NOTE When using chemical disinfection, carry-over residue from the cleaning process can interact adversely with the disinfecting agent. |  | - |
| 6.6.2 | Manual disinfection |  | P |
|  | If a manual disinfection method is specified, the following information shall be included, where applicable: | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | a) a description of the manual method with step-by-step instructions and the sequence of each individual process step; | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | b) a description of the process and processing parameters and, if applicable, any limits the medical device can withstand; | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | c) a description of the accessories required for the disinfection process; | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | d) identification and concentration of any process chemicals required for the disinfection process; | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | e) the contact time of any disinfecting agent used; | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | f) the quality of water to be used; | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | g) techniques to be used for rinsing; | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | h) if known, identification of any incompatibilities of disinfecting agents with the medical device. | See: [IFU] - Azurion R3.0 IFU, section 15.2.5 - Azurion R3.1 IFU, section 15.2.5 - Azurion R4.0 IFU, section 15.2.5 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | The medical device manufacturer's instructions for use may direct the processor to refer to the disinfectant manufacturer's instructions for use. | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
| 6.6.3 | Automated disinfection |  | N/A |
| 6.6.3.1 | If the automated disinfection process recommends the use of a washer-disinfector meeting the requirements of the ISO 15883 series, the information regarding the automated process may be limited to those parameters that are specific for the medical device, such as specific load configuration, positioning, connection, accessories, chemical (in the case of chemical disinfection), water quality, pressures or temperature limit(s) and a statement confirming the recommendation to use a washerdisinfector conforming with the ISO 15883 series. | Automated disinfection is not possible | N/A |
| 6.6.3.2 | If the specific disinfection requirements of the medical device do not allow a generic claim of compatibility with a washer-disinfector meeting the requirements of the ISO 15883 series, then the following information shall be included, where applicable: | Automated disinfection is not possible | N/A |
|  | a) a description of the process and processing parameters, including any limits the medical device can withstand; | Automated disinfection is not possible | N/A |
|  | b) a description of the accessories required for the disinfection process; | Automated disinfection is not possible | N/A |
|  | c) identification and concentration of any process chemicals required for the disinfection process; | Automated disinfection is not possible | N/A |
|  | d) the contact time of any disinfecting agent used; | Automated disinfection is not possible | N/A |
|  | e) the quality of water to be used; | Automated disinfection is not possible | N/A |
|  | f) techniques to be used for rinsing; | Automated disinfection is not possible | N/A |
|  | g) if known, identification of any incompatibilities of disinfecting agents with the medical device. | Automated disinfection is not possible | N/A |
|  | The medical device manufacturer's instructions for use may direct the processor to refer to the disinfectant manufacturer's instructions for use. | Automated disinfection is not possible | N/A |
|  | NOTE 2 Additional information conforming with 6.6.3.2 can be provided when the requirements of 6.6.3.1 are met if the medical device manufacturer chooses to do so. |  | - |
| 6.7 | Drying |  | P |
|  | Where drying is necessary, at least one drying method shall be specified. If a drying method is specified, the following information shall be included, where applicable: | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | a) a description of the process and processing parameters and, if applicable, any limits the medical device can withstand; | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | b) a description of the accessories required for the drying process; | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | c) specifications of the drying agent to be used; | No drying agent required. See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | N/A |
|  | d) the methods to be used and any special requirements to facilitate drying. | See: [IFU] - Azurion R3.0 IFU, section 15.2.4 - Azurion R3.1 IFU, section 15.2.4 - Azurion R4.0 IFU, section 15.2.4 [IFU- LegSupportBoard], section 4.2.1 [IFU- Foot Switch], section 5.1 [IFU- WFS], section 6.1 | P |
|  | NOTE Drying can be achieved as part of an automated cleaning and disinfection process. |  | - |
| 6.8 | Inspection and maintenance |  | P |
|  | Relevant information shall be provided if inspection, functionality testing, maintenance (including replacement of parts) or calibration of a medical device is required during or after processing to ensure proper function and safe use of that medical device. The following information shall be included, where applicable: | Inspection of parts is applicable, calibration or lubrication is not applicable. See: [IFU] - Azurion R3.0 IFU, section 15.2.6 - Azurion R3.1 IFU, section 15.2.6 - Azurion R4.0 IFU, section 15.2.6 [IFU- LegSupportBoard], section 4.1 [IFU- Foot Switch], section 5.2 [IFU- WFS], section 3.1.1 | P |
|  | a) the method(s) and performance criteria for inspecting the device with particular attention to medical device functionality including its impact on patient safety and safe use; | See: [IFU] - Azurion R3.0 IFU, section 15.2.6 - Azurion R3.1 IFU, section 15.2.6 - Azurion R4.0 IFU, section 15.2.6 [IFU- LegSupportBoard], section 4.1 [IFU- Foot Switch], section 5.2 [IFU- WFS], section 3.1.1 | P |
|  | b) the method to be used for adjustment and/or calibration of the medical device; | No adjustment or calibration is required. | N/A |
|  | c) the type, amount and method of application of lubricant; | No lubrication required during or after processing. | N/A |
|  | d) the instructions for reassembly of the medical device; | Disassembly and reassembly is not required. | N/A |
|  | e) a description of special tools to be used to maintain the medical device. | No special tools required for maintenance | N/A |
| 6.9 | Packaging |  | N/A |
|  | If a method for packaging and containing the medical device during and/or after processing is required, it shall be stated and be compatible with: | Packaging of the medical device during and/or after processing is not required | N/A |
|  | a) the specific conditions identified in the other processing stages; | Packaging of the medical device during and/or after processing is not required | N/A |
|  | b) the medical device. | Packaging of the medical device during and/or after processing is not required | N/A |
|  | c) environmental conditions during storage, transport or both. | Packaging of the medical device during and/or after processing is not required | N/A |
| 6.10 | Storage |  | P |
|  | Where applicable, information shall be provided on any specific limitations on time or conditions of storage of the processed medical device prior to use. | See: [IFU] - Azurion R3.0 IFU, section 17.5.1 - Azurion R3.1 IFU, section 17.5.1 - Azurion R4.0 IFU, section 17.5.1 [IFU- LegSupportBoard], section 5.2.1 [IFU- Foot Switch], section 5.4 [IFU- WFS], section 6.1 | P |
| 6.11 | Transportation |  | N/A |
|  | Where applicable, information shall be provided on any special requirements for the movement of a medical device from one location to the other. | Fixed equipment | N/A |
|  | NOTE 1 To prevent damage to the medical device during transport, the use of specific carts, containers, covers and other accessories can be recommended by the manufacturer. |  | - |
|  | NOTE2 Annex A contains further information regarding transportation to the point of use and offsite. |  | - |
| 7 | Presentation of the information |  | P |
|  | Processing instructions shall be provided. If these are available in electronic format, then printed format versions shall be available on request. Processing instructions shall contain the information required by Clause 6 as appropriate. | IFU is available in electronic format on the system, but can be downloaded from the internet or ordered as paper copy as described in the Leaflet [IFU-Leaflet] IFU addendum [IFU- LegSupportBoard] is provided as a hard copy with the Leg Support Board. | P |
|  | NOTE 1 An example format for giving detailed information for a particular medical device is given in Annex B. |  | - |
|  | NOTE 2 ISO 20417 provides guidance on the format and presentation of instructions for use. |  | - |